# LASER DENTAL INNOVATIONS

# StarLite Gold Fiber Optic

# User Instructions:

Sterilize before use unit comes not sterile.

# This is a  Precision Engineered Fiber Optic Assembly please read all instructions for best results

Protective

Cap

Fiber Strain

Relief

Laser Coupler

Using the Star Lite Fiber Optic:

Caution: Be careful to protect the front face of the optical ferrule at all times. Use protective cap when removed from the Laser. Distal end already has a tested good cleave for your convenience.

Cleaver Tip

- Sterilize the entire fiber optic assembly before use.   (See sterilization instructions)

- Remove from autoclave pouch.

- Remove protective ferrule cover.

- Carefully insert into the laser port and fully tighten.

- Strip and cleave distal end of fiber and insert thru your handpiece and tip.

Following the below procedure will result in a sterilization assurance level of 10-6 as verified by Northview Labs of Illinois

- Lightly wipe front ferrule surface with a lint free non abrasive cloth and IPA.

- Add protective front cap. If cap is lost be careful of front optical surface.

- Put in a large autoclave pouch that will allow maximum fiber coiling diameter. Coil diameter must be at least  6 inches.

- Avoid Ultrasonic Cleaning which may damage the front optical surface

- Autoclave at 270 F for 30 minutes

- Cool for 20 minutes before handling