StarLite Gold Fiber Optic

User Instructions

Sterilize before use unit comes not sterile.
This is a Precision Engineered Fiber Optic Assembly please read all instructions for best results

Using the StarLite Fiber Optic:
Caution: Be careful to protect the front face of the optical ferrule at all times. Use protective cap when removed from the Laser. Distal end already has a tested good cleave for your convenience.
1.	Remove from protective wrap or pouch.
2.	Wipe silicone fiber jacket with reagent grade IPA
3.	Sterilize the entire fiber optic assembly before use.   (See sterilization procedure)
4.	Remove protective ferrule cover.
5.	Carefully insert into the laser port and fully tighten.
6.	Remove tubing and cleave distal end of fiber and insert thru your handpiece and cannula.
Sterilization Procedure
1.	Wipe metal shell and rubber boot  with a  non abrasive cloth  and IPA             
2.	Add protective front cap. If cap is lost be careful of front optical surface.
3.   Cut off used glass fiber and remove 2 inches minimum of protective silicone tubing  
4.  Wipe exposed glass fiber with cold sterile solution make sure no debris or fluid remains in tubing
5.	Coil at least 8 inches in diameter
6.  Autoclave at 270 F (121 C) for 30 minutes
7.	Cool for 20 minutes before handling
8.   Using a sterile cleaver remove ½ inch of fiber from distal end
.-6

2956-006 A